Details of Bioreactor Diagram

View 1: Negative Pressure System & Chamber

- First Lung Chamber

  - Rectangular plexiglass chamber

  - Holds half of lung

  - 20” x 6” x 5”

- Second Lung Chamber

  - Will hold the second half of the lung

  - Will be a duplicate of A) and C) and D) – all specs the same

- First Cover (Bag Cover)

  - Covers the Lung Chamber

  - Bag with zip tie to keep vacuum

  - Bag allows for slight manipulations of lung without breaking seal

  - Can be clamped down onto base of chamber and has rubber to seal

  - (need to confirm that vacuum can be maintained through moveable bag and the zip tie)

- Second Cover (Hard Top Cover)

  - Clamps down to cover both the base of chamber and the Bag Cover

  - Secures the full chamber with additional rubber seals

- Three Peristaltic Pumps (10" x 8.5" x 8.5")

  - Each pump contains two pump heads

  - To be used to create negative pressure in the chamber

    - Each chamber will receive three lines (out of the six) from the pumps

  - Each pump head can deliver ~ 2900 mL/min

- Power Strip

  - Will allow power to full unit

    - All pumps

    - Computer system

- Computer System

  - Will be connected to pumps using included software to program pumps

  - (potentially also connected to all sensors)

View 2: Tube Connections to Lung (This will be duplicated for each Lung Chamber)

- Pulmonary Artery

  - Sewn around tube

  - Blood (or SDS/saline) goes in through this artery

- Pulmonary Vein

  - Sewn into tube

  - Blood (or SDS/saline) exits through this vein

- SDS/Saline Drain

  - After the SDS and saline are run through the lung the ‘tube switcher’ will dispose of these solutions

- Peristaltic Pump

  - (there will be one or two)

  - Will pump blood substitute from the reservoir through the Pulmonary Artery

- SDS Supply

  - Will be heated to 37 C

- Saline Solution Supply

  - Will be heated to 37 C

- Blood Substitute Reservoir

  - Will allow a changing of fluids approximately once every two days

  - Will be heated to 37 C

- Blood Backup

  - Used to replace blood every two days

  - Will have a manual switch to open when blood needs to be replaced that will go straight to the pump line

- Blood Disposal

  - Manual switch will allow ‘old’ blood to drain out

- Tube Switcher

  - After fluids come out of Pulmonary Vein the SDS and Saline will be sent to a disposal and the blood would go back to heated reservoir

- Tube Switcher

  - Can be switched from SDS process to Saline cleanout to Blood Substitute

- Trachea

  - Connected to tubes

  - SDS then Saline Solution then Perfusate will go through

- Bubble Releasers

  - Some sort of device used to release bubbles from all tubings

More Info

- Sensors

  - Heat Sensor

    - Confirm heat in chamber

    - (Checking blood and perfusate temperatures on the inbound tubes)

  - pH Sensor

    - In chamber to note break in sterility

    - On the outbound lines to note SDS remnants in the Saline Cleaning Solution

      - Advise of burst alveolus

  - Flow Sensor

    - On all inbound tubes to ensure that proper amounts are flowing